## Slide 1: Method Verification

## Slide 2: Introduction
An assessment focusing on how the analytical test procedure is suitable for its intended use under actual experimental condition
It is important in ensuring that measurement results reported to customers are correct

## Slide 3
It is done by
Calibration
Trueness/Accuracy
Precision
Linearity
Limit of detection
Limit of Quantification
Specificity

## Slide 4: Calibration
Calibration is the process of configuring an instrument to provide a result for a sample within an acceptable range
Nothing but relationship B/W the Instrument (Y) and measurement of Analyte (X)
Y=f(X)
Relationship is established by measurement of sample with known quantities of analyte

## Slide 5
Calibration function may be linear or curved in case of Immunoassay it may take special form – logistic in log x used in Immunoassay techniques and written in several forms.
Alternative model free approach is used to estimate smoothed spline curve often performed for Immunoassay
DISADVANTAGE: insensitive to aberrant calibration values

## Slide 6
If assumed calibration function does not reflect the true relationship a systematic error is associated with analytical method.
Common problem in Immunoassay is Hook Effect i.e deviated from expected calibration algorithm in high concentration range
If calibration function is linear and imprecision of signal response is same over the analytical measurement range.
Imprecision value increase the proportionally to the signal response then the analytical SD of the methods tends to increase proportionally to concentration

## Slide 7: Trueness
Closeness of agreement B/W the average value obtained from large series of results of measurements and true value
Difference B/W average value and true value is bias and expressed numerically inversely related to trueness
Expressed as Low, Medium, High
ISO- introduced trueness expression as the replacement for the term accuracy

## Slide 8: Accuracy
Closeness agreement B/W the result of measurement and true concentration of analyte and it is influenced by both Bias and imprecision this reflects total error
In relation to trueness, the concept
Recovery
Drift
carryover

## Slide 9: Recovery
Fraction by % increase in concentration measured in relation to the amount added
Carried out in the field of drug analysis
Recovery close to 100% is pre-requisite for high degree trueness

## Slide 10: Drift
Caused by the instrument or reagent instability over time so that calibration is biased

## Slide 11: Carryover
Assay must be close to ensure unbiased results
i.e placing the sample with known low value after a pathological sample with high value and observe
Increase can be stated as % of high value

## Slide 12: Precision
Closeness of agreement B/W independent results of measurements obtained under stipulated conditions
Degree of precision expressed on the bias of statistical measures of imprecision
SD/CV=SD/x
Imprecision solely related to randam error of measurement
Repeatability
Reproducibility

## Slide 13: Repeatability
Closeness of agreement B/W results of successive measurement carried out under the same condition

## Slide 14: Reproducibility
Closeness of agreement B/W results of measurement performed under changed condition of measurement
Two specifications often used
Total / B/W run precision – intermediate precision
Interlab precision

## Slide 15
Total SD
Within run components
Between run components

## Slide 16: Precision profile
Depends on the concentration of analyte being considered
Plotted in terms of SD or CV as a function of analyte concentration

## Slide 17: Linearity
Relationship B/W measurement and expected values over the analytical measurement range
Relative analyte concentration dilution sample may be used
Dilution carried out with appropriate sample matrix/ verified sample diluent
Evaluation conducted in various ways
More formal carried out on the bias of statistics

## Slide 18: Analytical measurement range
Analyte concentration, range over which measurements are within the declared tolerance for imprecision and bias of the method
Measurement of range extend from LLOQ to ULOQ for which performance specification are fulfilled

## Slide 19: Limit of detection
Characteristics of an assay defined as lowest value that significantly exceeds the measurement of a blank sample
Limit detected/estimated on the basis of repeated measurements of blank sample reported as mean +2SD., +3SD

## Slide 20: Analytical sensitivity
LLOQ not confused with analytical sensitivity
Ability of an analytical method to assess small difference in concentration of analyte

## Slide 21: Analytical specificity
Ability of an assay procedure to determine the concentration of target analyte without influence / factor

## Slide 22
Thank you